PFGE gels

PFGE gels are almost invariably run using genomic DNA embedded in agarose – see separate protocol for embedding yeast and mammalian cells.

Here I simply list a few of the common systems that we use for running the gels. It is really hard to optimise the run parameters, so we generally look for an existing PFGE programme with separation in the appropriate range (look at NEB and BioRad sites which give plenty of programs designed to separate their different ladders).

The only tricky thing about these gels is getting the agarose plugs into the wells, and everyone makes a mess of this the first few times. Otherwise these are just gels that run for a very long time. Take care of them though, and make the buffer fresh every time for best results.

We use BioRad Pulsed Field certified agarose (162-0137), but suspect that normal agarose would be just as good. However, we have learnt to avoid BioRad MegaBase agarose (which we previously recommended) – good samples often fail to resolve in this agarose for reasons that we have not been able to identify.

Digesting DNA in agarose plugs (optional)

NEB maintains a list of which enzymes work for digestion in agarose and which do not https://international.neb.com/tools-and-resources/usage-guidelines/digestion-of-agarose-embedded-dna-info-for-specific-enzymes

Cut plug in half and place in a 2ml tube (always 2ml tubes!) with 100µl 1x restriction buffer (normally CutSmart), incubate 15 min at room temperature

Replace the buffer with 100µl fresh buffer for another 15 min at room temperature

Replace the buffer with 100µl fresh buffer containing 10-20U restriction enzyme, incubate the plugs 4 hours - overnight at 37°

Equilibrate plugs 15 min in 1ml running buffer before loading onto gel.

Running, staining and blotting a PFGE gel

These conditions are for yeast chromosome XII (1 – 4Mb resolution) but the steps are always the same

Make 2.5L fresh 1x TBE, stir until fully dissolved

Microwave 120ml 1x TBE with 0.8% agarose until dissolved

For high molecular weight undigested yeast chromosomes use Biorad PFGE agarose instead (162-0137). For smaller fragments (up to 100kb or so), we just use normal multipurpose agarose (currently from Melford MB1200)

Stir until cooled to 55

Meanwhile set up gel caster and use a spirit level to make sure it is flat

Put the remaining 1x TBE in the tank, start pump, and set chiller unit to 14

Pour gel and let it set for >30min

Remove comb and fill wells with running buffer

Remove plug from wash buffer with tweezers

Put in Petri dish and cut in half (keep one half in wash buffer at 4)

Push ½ plug into well with tweezers

Ensure there are no air bubbles between plug and front of well

Remove gel caster edges, but leave gel on the black tray (don’t dislodge it!)

Put gel and tray in the fixture in the centre of the tank

Ensure that the wells are at the right end

Set the conditions and start the gel:

Voltage: 3V/cm

Run time: 68 hours

Switch time: 300-900s

Staining

When finished, move gel into 200ml water with 3l ethidium bromide,

Wash the tank with 2.5L water

After 30min, rinse the gel with water twice and take a picture

Blotting

Depurinate in fresh 0.25N HCl for 30 minutes (not more, not less!!)

Denature in 0.5N NaOH for 30 minutes

Neutralise twice for 15min with 0.5M Tris pH7.5, 1.5M NaCl

Transfer to HyBond N+ membrane in 6xSSC

PFGE for standard yeast chromosome (200kb – 1.5Mb resolution)

Gel percentage: 1.0% agarose

Running buffer: 0.5x TBE

Switch time: 60-120s

Run time: 24 hours

Voltage: 6V/cm

PFGE for meiotic yeast chromosomes (40 - 400kb resolution)

Gel percentage: 1.3% agarose

Running buffer: 0.5x TBE

Switch time: 15-25s

Run time: 24 hours

Voltage: 6V/cm

The following conditions are from Stults et al PMC2134781, they work in our hands

PFGE for human 5S rDNA (50 – 500kb resolution)

Digest plugs with EcoRV-HF

Gel:	1% agarose

0.5x TBE (fresh)

14°C

6V/cm

120° included angle

3-90s switch

24 hour run time

PFGE for human or mouse 45S rDNA (500kb – 6Mb resolution)

Digest plugs with EcoRV-HF (human) or PmeI (mouse)

Gel:	0.8% agarose

1x TAE (fresh)

14°C

2V/cm

106° included angle

300-2400s switch

92 hour run time

Solutions

1x TBE (2.5L):	27g Tris

13.75g Boric acid

10ml 0.5M EDTA pH8 (do not substitute with solid)

0.5x TBE (2.5L):	13.5g Tris

6.9g Boric acid

5ml 0.5M EDTA pH8 (do not substitute with solid)

For enzymes that require 50° incubation, increase volume to 200µl, this is ok for overnight

Using TBE diluted from even fairly new 10x stocks has a major negative impact on gel quality

Use a fresh Petri dish each time you load a gel, or wash it well

See our genomic southern protocol for detailed method, but note the differences in treatment here

Concentrated HCl (37% aqueous solution) is 12N

You need a CHEF-DRIII for this